{"text_format":"3","workbook":"xls","n_sheets":2,"sheets":["技术要求","技术参数"],"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"技术要求","index":0,"n_rows":9,"n_cols":5,"rows_kept":9,"merged":[[0,0,0,4],[7,0,7,1],[7,2,7,4],[8,0,8,4]],"flags":["formulas_cached_only"]}
["采购项目清单：人机共存空气消毒器等医疗设备采购项目"]
[]
["序号","项目名称","数量或规模","资金预算（万元）","资金来源"]
[1,"人机共存空气消毒器",10,17,"自筹"]
[]
[]
["合计",null,10,17]
["备注"]
["技术要求及配置确认书\n一、项目名称、数量：人机共存空气消毒器等医疗设备采购项目。\n二、项目用途说明：为满足医院诊疗需求。\n三、配置规格、附件及零配件（包括专用工具）： \n                                                       \n四、详细技术要求、参数及产品资料等：附技术偏离表、评分标准、产品彩页、相关证明文件\n\n\n五、售后服务条件及交货日期（或工期）："]

{"sheet":"技术参数","index":1,"n_rows":23,"n_cols":3,"rows_kept":23,"merged":[[0,0,0,2]],"flags":["formulas_cached_only"]}
["人机共存空气消毒器技术参数"]
[1,"应用场所：适用与门诊大厅，急诊大厅，普通手术室、产房、血液病区、烧伤病区、保护性隔离病区、重症监护病区；消毒供应中心检查包装灭菌区和无菌物品存放区、重症透析中心；检查室、治疗室、感染性疾病诊室等。"]
[2,"外形：移动式"]
[3,"外型尺寸：（L×W×H）：≤ 600×485×1940mm3"]
[4,"消毒方法：静电吸附型。"]
[5,"最大适用体积：250m3"]
[6,"额定循环风量：2500 m3/h"]
[7,"噪声：≤66 dB(A)"]
[8,"负离子浓度：≥6×106 个/cm3"]
[9,"消毒时空气中臭氧量：≤0.1mg/m3"]
[10,"静电场强度：≥7300V"]
[11,"可吸入颗粒物(0.5μm）平均净化效率为93.34%，提供权威机构检测报告。","★"]
[12,"气雾室细菌的杀灭率均大于99.90%(100%)，提供权威机构检测报告。","★"]
[13,"设备持续工作1小时，对体积为220 m3室内空气中的自然菌消亡率均大于90%。（提供权威检测报告）","★"]
[14,"可在有人状态下进行连续动态消毒。"]
[15,"选用优质、大容量的颗粒状活性炭组成的复合净化板专利技术，可有效去除有机气体和医院药水等各种异味。提供专利证明文件。","★"]
[16,"采用组合式高压静电场专利技术，除尘除菌效率高好，蜂窝六壁式结构积尘量大。提供专利证明文件。","★"]
[17,"采用高静压低噪声离心风机，循环风量大，具有高、中、低三档可调风速。"]
[18,"操作简单，程控自动运行，具有预约开机和关机功能，具有6个预约自动开关机时段功能，工作稳定可靠，确保准确性。","★"]
[19,"负氧离子清新净化空气，促进人体新陈代谢。"]
[20,"液晶中文显示屏，远红外线遥控。"]
[21,"自动累时维护报警，过滤器自动清洗报警，负离子正常工作指示功能\n高压静电场故障指示功能"]
[22,"输入功率：≤480W"]